SAMPLE PROCEDURES – Cleaning and Descaling the STERIS® Reliance® Vision® Multi-Chamber Washer/Disinfector

Product Numbers:

105105	LIQUID DESCALER (5 GAL - BUCKET)

105108	LIQUID DESCALER (4 X 1 GAL-CASE)

This document contains sample procedures for Cleaning and Descaling the STERIS Reliance Vision Multi-Chamber Washer/Disinfector.  The procedures contained in this document are only intended to provide a foundation for developing specific policies and procedures for your facility, and it is the individual responsibility of the health care facility to ensure compliance with applicable laws, regulations, and standards. The health care facility should seek expert advice and consultation for guidance with compliance issues. STERIS Corporation makes no representation, express or implied, with respect to compliance with local or federal laws, regulations, or standards. STERIS Corporation shall not be responsible for any loss, injury, damage, or claim arising from use of this document or the sample policies and procedures contained herein.

Frequency:   The Reliance Vision Multi-Chamber Washer/Disinfector should be cleaned daily and descaled on a weekly basis as recommended by STERIS.

Qualifications: Health Care personnel properly trained and competent in the maintenance and operation of the Reliance® Vision® Multi-Chamber Washer/Disinfector, thereby minimizing the risk of injury and ensuring optimal operating performance.

Procedure:

Daily Cleaning

On a daily basis:

- Empty the unload table and remove debris screen.

- Wipe each roller using a cloth dampened with household cleaner.

- Thoroughly wipe off cleaner and dry using a soft, lint-free cloth.

- Wipe plastic sides of guides and corners using a cloth dampened with 70% isopropyl alcohol or household cleaner.

- Thoroughly wipe off cleaner and dry using a soft, lint-free cloth.

- Rinse unload table and debris screen under running water.

- After the last cycle of the day, allow washer/disinfector to cool down and then rise and wipe all surfaces (inside and outside) of water outlet connection of wash and rinse chambers.

- Wipe the surface of the stainless port within the chamber.

Weekly Cleaning

- Clean the washer/disinfector exterior using a general purpose cleaner (nonabrasive) for general stains, a stainless steel cleaner to remove stubborn stains, and a stainless steel polish to keep equipment looking ike new. Proceed as follows:

  - Using a damp cloth or sponge, apply cleaner in a back-and-forth motion, rubbing in same direction as surface grain.

  - Thoroughly wipe off cleaner.

  - Polish surface with a clean, dry, lint-free cloth.

- Clean wash chamber interior with a moderately alkaline detergent solution. Rinse thoroughly with tap water and dry with a lint-free cloth.

Note: If interior is stained, use a general cleaner to remove general stains or a stainless steel stain remover for stubborn stains. Using a damp cloth or sponge, apply cleaner in a back-and-forth motion, rubbing in same direction as surface grain. Thoroughly wipe off cleaner and polish with a clean, dry, lint-free cloth.

- Clean wash chamber rotary spray arm assembly as follows:

  - Hold the fixed portion of the hub and unscrew rotary spray arm assembly.

  - Remove top rotary spray arm assembly from wash chamber ceiling.

  - Remove locking pin securing each spray arm on rotary spray arm hub. Remove spray arms.

  - Use fine wire (approximately wire gauge of paper clip) to clean sediment from spray nozzles. Rinse under running water to clean out sediment.

  - Reinstall spray arms on rotary spray arm hub.

  - Reinstall locking pin to secure both spray arms.

  - Reattach rotary spray arm assembly to top of wash chamber.

- Remove arms from bottom rotary spray arm assembly as follows:

  - Remove manifold sliding inlet.

  - Remove the three screws from spray arm hub. Remove hub and spray arm assembly.

- Clean spray arms using a fine wire and running water to clear out sediment.

- Reinstall spray arms as follows:

  - Insert spray arm assembly and hub over water inlet. Attach using the three screws previously removed.

  - Insert manifold sliding inlet on water inlet.

- Clean rotary spray arm assembly on washer racks as follows:

  - Pull on locking pin. Remove rotary spray arm from accessory.

  - Remove locking pins from both ends of spray arms. Pull on end plugs to remove.

  - Clean spray arms using a fine wire.

  - Rinse under running water to clear out sediment.

  - Reassemble spray arms and attach using locking pins.

  - Reinstall spray arm on washer racks.

Other Accessories

  - Take note of the position of the propelling jet oriented sideways, at the extremity of the spray arm.

  - Using a Phillips screwdriver, remove screw from rotary spray arm hub.

  - Holding the hub and the two gaskets, remove rotary spray arm from accessory rack.

  - Clean spray arms using a fine wire to push sediment inside spray arms.

  - Rinse under running water to clear out sediment through hub.

  - Reinstall spray arm on washer rack.

IMPORTANT: Rotary spray arm assembly must be reinstalled with propelling jet located to the right of the rack so the rotation is clockwise (the same rotation as the screw). If not, rotary spray arm assembly is upside down and will loosen with use.

- Clean door gasket using a soft cloth containing a mild detergent.

Weekly Descaling

On a weekly basis, decontaminate and descale washer/disinfector as follows, using DECONTAM cycle expressly designed for this purpose:

Note: The cycle consists of a 15 minute wash cycle at 150°F, followed by three hot water rinses (approximately 1 minute each).

- Wear appropriate Personal Protective Equipment (PPE) such as waterproof gloves and eye protection.

- Place empty two-level, three-level, four-level or five-level manifold rack on load table. NOTE: There is no bar code tag for the DECONTAM cycle.

- When chamber is ready, wash chamber door opens and manifold rack is indexed into wash chamber. Display shows:

- When wash chamber positioning switch detects rack is in proper position inside chamber, wash chamber load door closes. Display shows:

Note: On Select cycle to start screen, press Right arrow to scroll across touch screen to view following menus.

- Press DECONTAM on touch screen to select decontamination cycle. Display shows:

- DECONTAM cycle automatically starts.

- When DECONTAM cycle is completed, manifold rack is then indexed out of washer/disinfector and placed on unload table.

- Allow washer/disinfector to cool down and then open service access door of chambers.

- Check each debris screen for debris.

- Rinse each screen under running water if necessary. Rinse while it is still wet, before any debris dries.

- Reinstall debris screens in chambers.

- Close each chamber service access door.

IMPORTANT:  For each service access door, hold both handles while closing door. To help avoid leaks, both handles must be latched at the same time so door is properly closed.

Publication ID# M3975EN.2014-04

©2014 STERIS Corporation

All Rights Reserved.