SAMPLE PROCEDURES – Cleaning and Descaling the STERIS® Reliance® Vision® Single-Chamber Washer/Disinfector

Product Numbers:

105105	LIQUID DESCALER (5 GAL - BUCKET)

105108	LIQUID DESCALER (4 X 1 GAL-CASE)

This document contains sample procedures for Cleaning and Descaling the STERIS Reliance Vision Single-Chamber Washer/Disinfector.  The procedures contained in this document are only intended to provide a foundation for developing specific policies and procedures for your facility, and it is the individual responsibility of the health care facility to ensure compliance with applicable laws, regulations, and standards. The health care facility should seek expert advice and consultation for guidance with compliance issues. STERIS Corporation makes no representation, express or implied, with respect to compliance with local or federal laws, regulations, or standards. STERIS Corporation shall not be responsible for any loss, injury, damage, or claim arising from use of this document or the sample policies and procedures contained herein.

Frequency:   The Reliance Vision Single-Chamber Washer/Disinfector should be cleaned daily and descaled on a weekly basis as recommended by STERIS.

Qualifications: Health Care personnel properly trained and competent in the maintenance and operation of the Reliance Vision Single-Chamber Washer/Disinfector, thereby minimizing the risk of injury and ensuring optimal operating performance.

Procedure:

Daily Cleaning

On a daily basis:

- After the last cycle of the day, allow washer/disinfector to cool down, and then remove debris screen in bottom of sump and rinse under running water.

- Always clean debris screen while it is still wet, before debris dries.

Weekly Cleaning

- Once a week, allow washer/disinfector to cool down and then remove and wipe all surfaces (inside and outside) of water outlet connection in chamber.

- Inner surface of stainless port within the chamber must also be wiped.

Weekly Descaling

The following cleaning procedures should be performed once per week:

- Clean washer/disinfector exterior using a general purpose cleaner (nonabrasive) for general stains.  Use a stainless steel stain remover for stubborn stains and a stainless steel polish to keep equipment looking like new. Proceed as follows:

  - Using a damp cloth or sponge, apply cleaner in a back-and-forth motion, rubbing in the same direction as surface grain.

  - Thoroughly wipe off cleaner.

  - Polish surface with a clean, dry, lint-free cloth.

- Clean wash chamber interior with a moderately alkaline detergent solution.  Rinse thoroughly with tap water and dry with a lint-free cloth. If interior is stained, use a general purpose cleaner to remove general stains, or a stainless steel stain remover for stubborn stains. Proceed as follows:

  - Using a damp cloth or sponge, apply cleaner in a back-and-forth motion, rubbing in the same direction as surface grain.

  - Thoroughly wipe off cleaner.

  - Polish surface with a clean, dry, lint-free cloth.

On a weekly basis, decontaminate and descale washer/disinfector as follows, using DECONTAM cycle expressly designed for this purpose:

Note: The cycle consists of a 15 minute wash cycle at 150°F, followed by three hot water rinses (approximately 1 minute each).

- Wear appropriate Personal Protective Equipment (PPE) such as waterproof gloves and eye protection.

- Press OPEN DOOR on touch screen to open load door.

- Pour 414 ml (14 oz) of Liquid Descaler into chamber sump.

- Place empty two-level, three-level, four-level, or five-level Vision manifold rack in chamber.

NOTE: If SCS Conveyor System (option) is present, note that there is no bar code tag for the DECONTAM cycle.

- Press CLOSE DOOR on touch screen to close load door.  An audible signal sounds and display shows:

NOTE:  On Select cycle to start screen, press RIGHT arrow to scroll across touch screen until seeing DECONTAM cycle on touch screen.

- Press DECONTAM on touch screen to select decontamination cycle. Display shows:

- DECONTAM cycle automatically starts.

- When DECONTAM cycle is completed, an audible signal sounds and unload door opens. Display returns to Ready screen.

- Allow washer/disinfector to cool down.

- Bring a transfer cart next to washer/disinfector load door and ensure transfer cart brakes are applied (located on rear wheels of transfer cart). Ensure both transfer cart wheels are locked and swiveling is not possible.

- Carefully pull manifold rack onto transfer cart.

- Check each debris screen for debris.

- Rinse screen under running water if necessary. Rinse while it is still wet, before any debris dries.

- Reinstall debris screen.

- Press CLOSE DOOR on touch screen to close unload door.

Publication ID# M3979EN.2014-04

©2014 STERIS Corporation

All Rights Reserved.