SAMPLE PROCEDURES – Cleaning and Descaling the STERIS® Reliance® 1227 Cart and Utensil Washer/Disinfector

Product Numbers:

105105	LIQUID DESCALER (5 GAL - BUCKET)

105108	LIQUID DESCALER (4 X 1 GAL-CASE)

This document contains sample procedures for Cleaning and Descaling the STERIS Reliance 1227 Cart and Utensil Washer/Disinfector.  The procedures contained in this document are only intended to provide a foundation for developing specific policies and procedures for your facility, and it is the individual responsibility of the health care facility to ensure compliance with applicable laws, regulations, and standards. The health care facility should seek expert advice and consultation for guidance with compliance issues. STERIS Corporation makes no representation, express or implied, with respect to compliance with local or federal laws, regulations, or standards. STERIS Corporation shall not be responsible for any loss, injury, damage, or claim arising from use of this document or the sample policies and procedures contained herein.

Frequency:   The STERIS Reliance1227 Cart and Utensil Washer/Disinfector should be cleaned daily and descaled on a weekly basis as recommended by STERIS.

Qualifications: Health Care personnel properly trained and competent in the maintenance and operation of the STERIS Reliance 1227 Cart and Utensil Washer/Disinfector, thereby minimizing the risk of injury and ensuring optimal operating performance.

Procedure:

Daily Cleaning

After the last cycle of the day, clean debris screen located in bottom of washer/disinfector chamber as follows:

- Allow washer/disinfector to cool down before entering washer/disinfector chamber.

- Open chamber load door by turning door handle downward

- Use key to turn Lock/Unlock selector to Lock position to lock washer/disinfector.

Important: Retain key while inside chamber.

- Enter chamber.

- Remove debris screen located in bottom of chamber.

- Always clean debris screen while it is still wet, before debris dries. Brush debris screen to detach dirt.

- Rinse debris screen under running water.

- Reinstall debris screen in bottom of chamber. Press down firmly into place.

- Exit chamber.

- Use key to turn Lock/Unlock selector to Unlock position to unlock washer/disinfector.

- Close chamber load door and turn door handle upward.

Weekly Cleaning

- Clean washer/disinfector exterior using general purpose cleaner (nonabrasive) for general stains, stainless steel stain remover for stubborn stains and stainless steel polish to keep equipment looking like new. Proceed as follows:

  - Using damp cloth or sponge, apply cleaner in back-and-forth motion, rubbing in same direction as surface grain.

  - Thoroughly wipe off cleaner.

  - Polish surface with clean, dry, lint-free cloth.

- Open chamber load door by turning door handle downward.

- Before entering chamber to clean interior, use key to turn Lock/Unlock selector to Lock position to lock washer/disinfector.

IMPORTANT: Retain key while inside chamber.

- Enter chamber.

- Clean wash chamber as follows:

  - Using a damp cloth or sponge, apply a moderately alkaline detergent solution in a back-and-forth motion, rubbing in the same direction as surface grain.

  - If interior is stained, use a general cleaner to remove general stains or a stainless steel stain remover for stubborn stains.

  - Thoroughly wipe off cleaner and polish with a clean, dry, lint-free cloth.

  - Rinse thoroughly with tap water and dry with a lint-free cloth.

  - Polish surface with a clean, dry, lint-free cloth.

  - Once per month, clean each door gasket using a brush and a germicidal detergent.

- When chamber interior cleaning is completed, exit chamber.

- Use key to turn Lock/Unlock selector to Unlock position to unlock washer/disinfector.

- Close chamber load door and turn door handle upward.

Weekly Descaling

Decontaminate and descale chamber as follows, using the DECONTAM cycle expressly designed for this purpose:

- Press CYCLES on Ready screen. Display shows Select cycle to start screen:

- Press Right arrow to scroll across touch screen to view DECONTAM cycle button.

- Press Decontam button on touch screen. Display shows time countdown:

- After countdown has started, display shows:

- Follow instructions on touch screen to proceed with manual injection.

- When manual injection is completed, press Done button. Display shows time countdown again.

- If present, optional reservoir drains completely to ensure only clean water is being used.

- DECONTAM cycle proceeds with filling, injecting, recirculating and draining phases to decontaminate wash chamber and optional reservoir (if present).

NOTE:  DECONTAM cycle lasts about 45 minutes (without optional reservoir) or about 60 minutes (with optional reservoir).

- When cycle is completed, open unload door by turning door handle downward and allow chamber to cool down.

- Before entering chamber, use key to turn Lock/Unlock selector to Lock position to lock washer/disinfector.

IMPORTANT: Retain key while inside chamber.

- Enter chamber.

- Verify debris screen in bottom of chamber of debris. If required, remove and rinse off any debris.

- Reinstall debris screen at the bottom of chamber. Press down firmly to lock in place.

- Exit chamber.

- Use key to turn Lock/Unlock selector to Unlock position to unlock washer/disinfector.

- Close door and turn door handle upward.

- Display returns to Ready screen:

Publication ID# M3976EN.2014-04

©2014 STERIS Corporation

All Rights Reserved.