SAMPLE PROCEDURE

Routine Monitoring of STERRAD® Sterilization Systems

Product Numbers:

LCB020 Verify® V24 Self Contained Biological Indicator

LCB004 Verify® Self-Contained Biological Indicator HP Activator

This document contains components of a sample procedure for the routine monitoring of STERRAD Sterilization Systems. The procedures contained in this document are only intended to provide a foundation for developing specific policies and procedures for your facility. It is the responsibility of the healthcare facility to ensure compliance with applicable laws, regulations, standards, and industry recommended practices.  The healthcare facility should seek expert advice and consultation for guidance with compliance issues.  STERIS Corporation makes no representation, express or implied, with respect to this sample procedure’s compliance with local or federal laws, regulations, standards, or industry recommended practices.  STERIS Corporation shall not be responsible for any loss, injury, damage, or claim arising from use of this document or the sample policies and procedures contained herein.

Title: Routine Monitoring of STERRAD Sterilizers.

Definitions:

Mechanical Monitor:  Sterilizer time, temperature, and pressure recording devices.

Biological Indicator:  Test systems containing viable microorganisms providing a defined resistance to a specified sterilization process (ANSI/AAMI ST58:2013). A biological indicator does not verify that an item is actually sterile.

Biological Indicator “Control”: An unprocessed biological indicator used to ensure viable organisms are present in the indicator lot and to monitor the operation of the incubator.

Chemical Indicators: System that reveals a change in one or more predefined process parameters based on a chemical or physical change resulting from exposure to a process (ANSI/AAMI/ISO 11140-1:2005).

Policy:

Routine microbial challenges shall be conducted daily for each sterilizer listed below.

- STERRAD® 100S Sterilizer

- STERRAD® 200 Sterilizer

- STERRAD® NX Sterilizer

- STERRAD® 100 NX Sterilizer

Procedure:

Preparation and Placement:

- Load the sterilizer with items to be processed according to the sterilizer Operator Manual.

- Place one self-contained biological indicator (SCBI) into the (Insert name of sterilization pouch cleared for use with the STERRAD Sterilizer). Insert one (Insert name of CI cleared for use with the STERRAD sterilizer) into the pouch such that the indicator ink is visible through the clear side of the pouch. Ensure that the expiration date has not passed for the SCBI, CI or sterilization pouch.

- Place the pouch at the back of the bottom shelf per the sterilizer’s operator manual.

NOTE: Do not stack pouch directly on or beneath other items on the shelf.

- Close the chamber door.

- Initiate the appropriate sterilization cycle.

- Default cycle of the STERRAD® 100S Sterilizer

- Default cycle of the STERRAD® 200 Sterilizer

- Standard or Advanced Cycles of the STERRAD® NX Sterilizer

- Express, Flex Scope or Standard Cycles of the STERRAD® 100 NX Sterilizer

Upon completion of the cycle:

- Put on gloves.

- Review the cycle printout to ensure that cycle did not abort and follow directions for processing a sterilized load in the sterilizer User’s Guide.

  - If cancellation messages show parameters were not met, the sterilization cycle was not successful. The processed items may not be released for use.  Reject the cycle. Repackage and sterilize all items.

- Remove the pouch from the shelf. Look for any leakage of the growth media onto the pouch or in the Biological Indicator’s (BI) vial, if no leakage proceed to step 4.

  - Pouches which show leakage must be placed in a steam compatible container or pouch for disposal. Refer to the SCBI instructions for disposal recommendations.

- Remove the chemical indicator strip. Observe the strip for a “Pass” color change per the CI’s instructions for use.

  - If the chemical indicator shows passing results, continue to step 5.

  - If the chemical indicator shows failing results, follow departmental procedures for investigating suspected sterilization failures.

- Remove the SCBI and check the chemical indicator on the SCBI vial label for a color change from magenta to orange/yellow.

  - If the chemical indicator is orange/yellow, proceed to step 6.

  - If the chemical indicator is not orange/yellow, determine the root cause of the cycle failure. Follow departmental procedures for investigating suspected sterilization failures.

- Seal and activate the BI using the Verify Self Contained Bioloigcal Indicator HP Activator. The BI is properly actvated when the cap is pushed down to the second bar on the vial label and the culture media has been released from the crushed ampoule.

- Label the SCBI with pertinent process information.

- Place the activated SCBI in a Verify Incubator heated to 55 - 60°C. Incubate for 24 hours.

Control

A control must be performed once each day that a microbial test is performed, and whenever the lot of SCBIs changes.

- Obtain an SCBI from the same lot used in the microbial challenge load.

- Seal, activate, and incubate the SCBI as described in step 6 through 8 above.

Interpretation of Biological Results

- Check the incubating “Test” and “Control” SCBIs for growth at 24 hours (up to 7 days) of incubation.

  - If the liquid remains orange with no turbidity (cloudiness), the biological indicator is negative (no growth).

  - If the liquid turns from orange to yellow or turbidity (cloudiness) is seen, the biological indicator is positive (growth).

- Record the SCBI “Test” and “Control” results.

- The test passes when the “Test” SCBI is negative (no growth is seen) and the “Control” SCBI is positive (growth is seen).

- The test fails if the “Test” SCBI is positive (growth is seen) and the “Control” SCBI is positive (growth is seen).  Remove the sterilizer from service. Follow departmental procedures for investigating suspected sterilization failures.

- The test is inconclusive whenever the “Control” biological indicator is negative (no growth is seen).  The test is invalidated.  Determine the reason for the failure. Repeat the test.

# Footnotes

STERRAD is the registered trademark of Advanced Sterilization Products, Division of Ethicon Inc., a Johnson & Johnson company