SAMPLE PROCEDURE

Immediate Use Steam Sterilization (IUSS) Load Monitoring with VERIFY™ STEAM Integrating Indicator 5CM and VERIFY™ Dual Species Self-Contained Biological Indicators

Product Numbers:

PCC067	VERIFY™ STEAM Integrating Indicator 5CM

PCC067B 	VERIFY™ STEAM Integrating Indicator 5CM

S3060	VERIFY™ Dual Species Self-Contained Biological Indicator

S3061	VERIFY™ Dual Species Self-Contained Biological Indicator

This document contains sample procedures for biological and chemical monitoring of immediate use steam sterilization cycles. The procedures contained in this document are only intended to provide a foundation for developing specific policies and procedures for your facility. It is the responsibility of the health care facility to ensure compliance with applicable laws, regulations, standards, and industry-recommended practices. The health care facility should seek expert advice and consultation for guidance with compliance issues. STERIS Corporation makes no representation, express or implied, with respect to this sample procedure’s compliance with applicable laws, regulations, standards, and industry-recommended practices. STERIS Corporation shall not be responsible for any loss, injury, damage, or claim arising from use of this document or the sample policies and procedures contained in it.

Title: Immediate Use Steam Sterilization (IUSS) Load Monitoring with VERIFY™ STEAM Integrating Indicator 5CM and VERIFY™ Dual Species Self-Contained Biological Indicators.

Definitions:

Mechanical Monitor: Sterilizer time, temperature, and pressure recording devices.

Biological Indicator: Test systems containing viable microorganisms providing a defined resistance to a specified sterilization process (ANSI/AAMI ST79 A4:2013).

Biological Indicator “control”: An unprocessed biological indicator used to ensure viable organisms are present in the indicator lot and to monitor the operation of the incubator.

Chemical Indicators (CI): Systems that show exposure to sterilization processes by means of physical and/or chemical change of substances, and which are used to monitor the attainment of one or more of the variables required for sterilization process (ANSI/AAMI/ISO 11140-1).

Release of Immediate Use Steam sterilized items (excluding implants)

Procedure:

Table 1: List of Validated IUSS Cycles

| Sterilization Cycle Type | Sterilization Temperature and Time |
| --- | --- |
| Immediate Use Steam Sterilization | 270 °F / 132 °C for 3 and 10 minute Gravity |
|  | 270 °F / 132 °C for 4 minute Dynamic Air Removal |
|  | 275 °F / 135 °C for 3 and 10 minute Gravity |
|  | 275 °F / 135 °C for 3 minute Dynamic Air Removal |

Initiating the cycle:

- Remove one VERIFY STEAM Integrating Indicator 5CM from the pouch. Ensure that it is within its expiration dating.

- Place the integrating indicator in the tray or container next to the item being IUSS sterilized. Place the tray or container into the steam sterilizer. (Insert IUSS sterilization container instructions for use here.)

- Close the chamber door.

- Process the load using a validated immediate use steam sterilization cycle (IUSS). See Table 1.

Upon completion of the cycle:

- Review the cycle printout to ensure that processing parameters were met (i.e., time, temperature and preconditioning pulses).

- Open the chamber door and carefully remove the tray or container that was IUSS sterilized.

- Aseptically remove the integrator strip from the tray or container and read the results.

- Examine the window on the indicator strip.

- The strip indicates a “Pass” if the dark bar has reached the ACCEPT area. The items may be used.

- The strip indicates a “Fail” if the dark bar has not reached the ACCEPT window. The items may NOT be used. Follow departmental procedures for investigating suspected sterilization failures.

- File the cycle printout and integrating indicator in accordance with departmental documentation procedures.

Release of immediate use steam sterilized implantable devices

IMPLANTABLE DEVICES SHOULD BE IUSS STERILIZED ONLY WHEN A DOCUMENTED 
EMERGENCY NEED HAS OCCURRED

NOTE: AORN and AAMI do not recommend immediate use steam sterilization of implantable devices unless a documented emergency necessitates the process.

Procedure:

Initiating the cycle:

- Identify the implant that must be immediate use steam sterilized and complete an Exception Form for Premature Release of an Implantable Device.

- Complete the early release log form.

- Ensure that the integrating indicator and self-contained biological indicator (SCBI) is within expiration.

- Place one integrating indicator and one VERIFY Dual Species Self-Contained Biological Indicator in the tray or container next to the item being sterilized. Place the tray or container into the steam sterilizer. (Insert immediate use steam sterilization container instructions for use here.)

- Close the chamber door.

- Process the load using a validated immediate use steam sterilization cycle (IUSS). See Table 1

Upon completion of the cycle:

- Review the cycle printout to ensure that processing parameters were met (i.e., time, temperature preconditioning pulses).

- Open the chamber door and carefully remove the tray or container that was immediate use steam sterilized.

- Aseptically remove the integrator strip from the tray or container and read the results.

- Examine the window on the indicator strip.

- The strip indicates a “Pass” if the dark bar has reached the ACCEPT area. Proceed to step 10.

- The strip indicates a “Fail” if the dark bar is not reached the ACCEPT area. The items may NOT be used. Follow departmental procedures for investigating suspected sterilization failures.

- Complete the Exception Form for Premature Release of an Implantable Device and file with the early release log.

- File the cycle printout and integrating indicator in accordance with departmental documentation procedures.

- Check the chemical indicator on the self-contained biological indicator (SCBI) vial label for a color change from blue to brown or black.

  - If the indicator is not brown or black, the immediate use steam sterilized items may NOT be used. Follow departmental procedures for reporting sterilization failures.

  - If the indicator is brown or black, and the vial is cool to the touch, firmly seal the vial using the activator and cap for the VERIFY Dual Species SCBI. Place the vial into the activator base. Cover the vial with the activator cap. Place the activator set on a hard surface. Push down on the activator cap until the SCBI cap is aligned with the second black bar on vial label.

- To activate, remove the SCBI from the activator set push or pull the SCBI completely through the restricted space on the activator.

- Label the SCBI with pertinent process information.

- Place the activated SCBI in the VERIFYTM Incubator heated to 55 - 59°C.

- A control must be performed once each day that a SCBI is used.

  - Obtain an SCBI from the same lot as that used in the test.

  - Seal, activate and incubate the SCBI as described in steps 13 through 15.

- Check “Test” and “Control” SCBIs each shift and at 24 hours of incubation for growth.

  - If the liquid remains blue and with no turbidity (cloudiness), the biological indicator is sterile (no growth). This indicates that conditions were adequate to achieve sterilization.

  - If the liquid turns from blue to yellow or turbidity (cloudiness) is seen, the biological indicator is not sterile (growth). This indicates that conditions were inadequate to achieve sterilization.

  - Record the SCBI “Test” and “Control” results.

- The biological kill is confirmed when the “Test” SCBI is sterile (no growth is seen) and the “Control” SCBI is not sterile (growth is seen).

- The biological kill is NOT confirmed when the “Test” SCBI is not sterile (growth is seen) and the “Control” SCBI is not sterile (growth is seen). Follow departmental procedures for reporting sterilization failures and patient follow-up.

- The biological kill test is inconclusive whenever the “Control” biological indicator is sterile (no growth is seen). The test is invalidated.

- Complete the Exception Form when the biological results are known and file with the cycle records.